## Eingaben
Settings
| Basic Type | Pressure Rating | Display | Accuracy
DFC | [5 empty] | Min recommended diameter is 25mm, please use Venturi
| enter Inside Diameter |  |  | 107.3 | mm
Gas type |  |  | Density [kg/Nm³] | 1.2929 |  | 1.23
Normal Operation Pressure [barabs] |  |  | 1 | Maximum 14bar |  | FALSE
Normal Operation Temperature [°C] |  |  | 20 | Maximum 180°C |  | FALSE
Max Operation Flow |  |  | 800
ImprooveIT |  |  | 1 | [-]
Damping
4..20mA Output |  |  | 4..20mA=-50..500°C
0..10VDC Output |  |  | 0..10VDC=0..14bar abs
Calculated Results
| kg/h |  | Nm³/h |  | Druckverlust [mbar]
Very Low Flow** | 565.4823 |  | 437.3751 |  | 0.28985 | 0 | FALSE
Low Flow | 1696.4468 |  | 1312.1253 |  | 2.6087 | 1 | TRUE
Medium Flow | 5654.8227 |  | 4373.751 |  | 23.4781 | 0 | FALSE
Maximum Flow | 16964.468 |  | 13121.253 |  | 211.3033 | 0 | FALSE
Operational Flow |  |  | 800 | Nm³/h
Operational Pressure |  |  | 1 | bar abs
Operational Temperature |  |  | 20 | °C
Operational Density |  |  | 1.1889 | kg/m³ |  | Fluid Velocity | 26.7 | m/s
dp at Operational Flow |  |  | 13.9632 | mbar
Pressure Loss at Operational Flow |  |  | 0.78548 | mbar
Epsilon |  |  | 0.99601 | [-]
KEPS |  |  | 1065.2754 | [-]
DIP switches settings |  |  |  |  | Display Settings (Option D1)
1 | ON | Flow Span Low |  |  | FlowUnit | Nm³ |  | TRUE
2 | OFF | Flow Span High |  |  | TimeBasis | h |  | FALSE
3 | ON | Damping Low |  |  | FlowSpan | 13121.253 |  | TRUE
4 | OFF | Damping High |  |  | PressRange | 14bar (HP) |  | FALSE
5 | ON | 4..20mA Output |  |  | KEPS | 1065 |  | TRUE
6 | ON | 0..10VDC Output |  |  | Kspan | 750 |  | TRUE
7 | OFF | DIAGNOSIS | [5 empty] | FALSE
Accuracy (typical)
* Please proceed zero offset after installation and in an adequate service intaervall
V1.07 OB 10.04.2019
© systec Controls Mess- und Regeltechnik GmbH, Lindberghstraße 4, 82178 Puchheim, Germany
www.systec-controls.de - Fon: +49-(0)89-809060 - Fax: +49-(0)89-80906-200 - Email: info@systec-controls.de
Note on C4: Select Basic Type in accordance with Type number (See Type Label): P (S/E/Q)-Insertion Probe Type V40-1 1/2" Venturi V25-1 " Venturi V20-3/4" Venturi 10mm core V8-3/4" Venturi 8mm core V6-3/4" Venturi 6mm core V4- 3/4" Venturi 4mm core
Note on D4: LP - Low Pressure Type, 0..4,5 bara HP - high Pressure Type, 0..14 bara
Note on E4: D0 - No Display D1 - Display (std) F12 - Remote Display
Note on F4: DS - Standard Accuracy (std) DH - High Accuracy
Note on B9: Please enter the normal operation pressure of your application
Note on B10: Enter the normal operation temperature of your application
Note on B11: Enter the maximum flow of your application
Note on B12: ImproovIT is a calibration factor for the flowmeter: q corr = q std * ImprooveIT
Note on B13: Damping of the analouge flow output
Note on B14: 4..20mA Output function
Note on B15: 0..10VDC output function

## Berechnungen (hidden)
Typeb |  | Pressure |  | Accuracy |  | Display |  | Gas Type |  | User Units |  | Damping |  | Ausgang 4..20 |  | Ausgang 0..10VDC
P10 (C/S/Q) | 1 | HP | 1 | DH (opt) | 2 | D1 | 1 | Air | 1 | kg/h | 4 | 0,2 sec | 3 | Flow | 2 | Flow | 2
V40 |  | LP |  | DS |  | D0 (opt) |  | Nitrogene |  | kg/min |  | 1 sec |  | Temperature |  | Pressure
V25 | [5 empty] | F12 (opt) |  | Methane |  | kg/s |  | 3 sec
V20(10) | [7 empty] | Carbon dioxide |  | Nm³/h |  | 10 sec
V16(8) | [7 empty] | Oxygene |  | Nm³/min
V16(6) | [7 empty] | Ammonia |  | Nm³/s
V16(4) | [7 empty] | other… |  | m³/h [compressor inlet, 20°C, 1,0bar]
m³/min [compressor inlet, 20°C, 1,0bar]
m³/s [compressor inlet, 20°C, 1,0bar]
Nl/min
Messbereiche Luft von Venturis
Span [kg/h] air | HP | LP
V40 | 975.5767 | 553.1
V25 | 603.7799 | 342.311
V20 | 260.1655 | 147.5
V8 | 163.7005 | 92.81
V6 | 83.0426 | 47.081
V4 | 36.5747 | 20.736 | 02.04.2019 Gemessen
Messbereiche Luft von Sonde, ermittelt als Polynom 5. Ordnung für 3 Durchmesserbereiche (Curve Fit)
Span P [kg/h] | [6 empty] | ID² | 11513.29
diameter | ID² | HP |  | 5th Degree Polynomial Fit: y=a+bx+cx^2+dx^3...
25 | 625 | 561 |  | Coefficient Data: |  |  | ID0..50
30 | 900 | 1124 |  | a = | -2.9224
35 | 1225 | 1535 |  | b = | -0.26861 |  | -5496731.4528 | Flow [kg/h]
40 | 1600 | 2036 |  | c = | 0.003139
45 | 2025 | 2626 |  | d = | -2.2636e-06
50 | 2500 | 3305 |  | e = | 6.8419e-10
65 | 4225 | 5848 |  | f = | -7.1562e-14
75 | 5625 | 7975
88 | 7744 | 11175 |  | 5th Degree Polynomial Fit: y=a+bx+cx^2+dx^3...
100 | 10000 | 14509 |  | Coefficient Data: |  |  | ID50..200
200 | 40000 | 64066 |  | a = | -69.9803
300 | 90000 | 130323 |  | b = | 1.3435 |  | 16964.468
500 | 250000 | 375055 |  | c = | 1.4531e-05
1000 | 1000000 | 1565268 |  | d = | -2.5076e-10
1500 | 2250000 | 3548957 |  | e = | 1.3371e-15
1750 | 3062500 | 4838354 |  | f = | -2.2261e-21
2000 | 4000000 | 6326121
5th Degree Polynomial Fit: y=a+bx+cx^2+dx^3...
|  |  |  | Coefficient Data: |  |  | ID 200..unendlich
|  |  |  | a = | 1581.9956
|  |  |  | b = | 1.4535 |  | 18343.438
|  |  |  | c = | 2.0578e-07
|  |  |  | d = | -1.2557e-13
|  |  |  | e = | 3.2978e-20
|  |  |  | f = | -3.1129e-27
|  |  |  | Flow HP | Flow LP
|  |  |  | 16964.468 | 9617.9493
Auswahl Primärelementyp
| HP Flow | LP Flow | kg/h für Luft
| 16964.468 | 9617.9493
Gaskorrektur für andere Gase als Luft |  |  | kg/h für sonstiges Gas
| HP Flow | LP Flow
| 16964.468 | 9617.9493
Faktor für Umrechnung auf Usereinheiten
|  |  |  | Normdichte | 1.2929 | Dichte 20°C/1bar | 1.1889
1 | kg/h | 1
2 | kg/min | 0.016667 |  | Umrechnungsfaktor | 0.77346
3 | kg/s | 0.00027778
4 | Nm³/h | 0.77346
5 | Nm³/min | 0.012891
6 | Nm³/s | 0.00021485
7 | m³/h [compressor inlet, 20°C, 1,0bar] | 0.84109
8 | m³/min [compressor inlet, 20°C, 1,0bar] | 0.014018
9 | m³/s [compressor inlet, 20°C, 1,0bar] | 0.00023363
10 | Nl/min | 12.8909
Berechnung von Druckverlusten und Expansionszahlen, für Luft bei Normbedingungen, qmax, dp=400mbar | [15 empty] | 25 | 82
[10 empty] | Polynomisch für Expansionszahl | [5 empty] | 30 | 69.2
Aus DF12 | ID | dp loss bei 200mbar | dploss bei 400mbar | Epsilon bei 200mbar | Epsilon bei 400mbar |  | Polynomisch für Epsilon bei 400mbar | Polynomisch Druckverlust bei 400mbar |  | Exponential Association: y=a(1-exp(-bx) | [5 empty] | 35 | 60
deltacalc | 25 | 41 | 82 | 0.848 | 0.696 |  | 0.7389 | 88.8316 |  | Coefficient Data: | [5 empty] | 40 | 53.2
| 30 | 34.6 | 69.2 | 0.897 | 0.794 |  | 0.78357 | 67.1059 |  | a = | 0.88756 |  |  |  |  | 45 | 47.8
| 35 | 30 | 60 | 0.918 | 0.836 |  | 0.81482 | 55.284 |  | b = | 0.071473 |  |  |  |  | 50 | 43.6
| 40 | 26.6 | 53.2 | 0.928 | 0.856 |  | 0.83668 | 47.7802 | [7 empty] | 65 | 34.8
| 45 | 23.9 | 47.8 | 0.933 | 0.866 |  | 0.85197 | 42.5559 | [7 empty] | 75 | 30.8
| 50 | 21.8 | 43.6 | 0.936 | 0.872 |  | 0.86266 | 38.6868 | [7 empty] | 88 | 27.02
| 65 | 17.4 | 34.8 | 0.94 | 0.88 |  | 0.87904 | 31.3283 |  | Polynomisch für Druckverlust | [5 empty] | 100 | 24.4
| 75 | 15.4 | 30.8 | 0.941 | 0.882 |  | 0.88339 | 28.2741 |  | Harris Model: y=1/(a+bx^c) | [5 empty] | 150 | 11.82
| 88 | 13.51 | 27.02 | 0.942 | 0.884 |  | 0.88592 | 25.4208 |  | Coefficient Data: | [5 empty] | 200 | 8.88
| 100 | 12.2 | 24.4 | 0.942 | 0.884 |  | 0.88686 | 23.4714 |  | a = | -0.087082 |  |  |  |  | 300 | 5.92
| 150 | 5.91 | 11.82 | 0.942 | 0.884 |  | 0.88754 | 18.6787 |  | b = | 0.051725 |  |  |  |  | 500 | 3.56
| 200 | 4.44 | 8.88 | 0.942 | 0.884 |  | 0.88756 | 16.1688 |  | c = | 0.1996 |  |  |  |  | 1000 | 1.76
| 300 | 2.96 | 5.92 | 0.942 | 0.884 |  | 0.88756 | 13.4406 | [7 empty] | 1500 | 1.18
| 500 | 1.78 | 3.56 | 0.942 | 0.884 |  | 0.88756 | 10.901 | [7 empty] | 1750 | 1
| 1000 | 0.88 | 1.76 | 0.942 | 0.884 |  | 0.88756 | 8.4554 | [7 empty] | 2000 | 0.8
| 1500 | 0.59 | 1.18 | 0.942 | 0.884 |  | 0.88756 | 7.3759
| 1750 | 0.5 | 1 | 0.942 | 0.884 |  | 0.88756 | 7.0158
| 2000 | 0.4 | 0.8 | 0.942 | 0.884 |  | 0.88756 | 6.7233
| Druckverlust sonde bei qmax Luft; normbedingungen |  |  |  | 22.5014 |  |  | 22.5014
|  | Innendurchmesser |  | 107.3
|  | Epsilon bei 400mbar |  | 0.88715
|  | Epsilon bei operation |  | 0.99601 | 0.99601
|  | dploss bei 100mbar | dploss bei 400mbar | Epsilon bei 100mbar | Epsilon bei 400mbar
ID 20/35 | V40 | 13.72 | 54.88 | 0.9381 | 0.7524
ID 25/15 | V25 | 13.24 | 52.96 | 0.9362 | 0.7448
ID 15/10 | V20 | 12.5 | 50 | 0.9301 | 0.7204
ID 8/15 | V18(8) | 13 | 52 | 0.932 | 0.728 | 932
ID 6/15 | V18(6) | 14 | 56 | 0.934 | 0.736
ID 4/15 | V18(4) | 16 | 64 | 0.936 | 0.744
|  | Epsilon | 0.88715
|  | Keps | 1065.2754
Berechungen für Operation conditions
Dichte bei 14 bar
| Normdichte | 1.2929 |  | 17.8639
| Betriebsdichte | 1.1889
Umrechnung User Operation flow auf kg/h |  |  |  | 1034.32
dp Berechnung |  | 13.9632 | mbar
druckverlust berechnung |  | 0.78548 | mbar
| Berechnungen für Druckverlust und Unsicherheit |  |  |  |  | Spanfehler | 0.02
| Messbereichsendwert |  | 1312.1253
| Messbereich | Flow | dp | Fehler 0.3mbar | Fehler 0,1 mbar | Accuracy o.s. [no offset] | Accuracy o.s. [after offset *] | Uncertaincy Std Accuracy | Uncertaincy High Accuracy | Uncertaincy o.s. Std, Accuracy | Uncertaincy o.s. High Accuracy
| 0.05 | 65.6063 | 0.093907 | 106.8% | 45.7% | 5.34% | 2.285% | 106.8% | 45.7% | 0.534% | 0.2285%
| 0.1 | 131.2125 | 0.37563 | 36.11% | 14.53% | 3.611% | 1.453% | 36.11% | 14.53% | 0.3611% | 0.1453%
| 0.15 | 196.8188 | 0.84516 | 18.4% | 7.751% | 2.76% | 1.163% | 18.4% | 7.751% | 0.276% | 0.1163%
| 0.2 | 262.4251 | 1.5025 | 11.53% | 5.274% | 2.306% | 1.055% | 11.53% | 5.274% | 0.2306% | 0.1055%
| 0.25 | 328.0313 | 2.3477 | 8.197% | 4.108% | 2.049% | 1.027% | 8.197% | 4.108% | 0.2049% | 0.1027%
| 0.3 | 393.6376 | 3.3806 | 6.343% | 3.468% | 1.903% | 1.04% | 6.343% | 3.468% | 0.1903% | 0.104%
| 0.35 | 459.2439 | 4.6014 | 5.208% | 3.081% | 1.823% | 1.078% | 5.208% | 3.081% | 0.1823% | 0.1078%
| 0.4 | 524.8501 | 6.01 | 4.465% | 2.829% | 1.786% | 1.131% | 4.465% | 2.829% | 0.1786% | 0.1131%
| 0.45 | 590.4564 | 7.6064 | 3.953% | 2.655% | 1.779% | 1.195% | 4% | 2.655% | 0.18% | 0.1195%
| 0.5 | 656.0626 | 9.3907 | 3.585% | 2.531% | 1.792% | 1.266% | 4% | 2.531% | 0.2% | 0.1266%
| 0.55 | 721.6689 | 11.3627 | 3.312% | 2.439% | 1.821% | 1.341% | 4% | 2.439% | 0.22% | 0.1341%
| 0.6 | 787.2752 | 13.5226 | 3.103% | 2.369% | 1.862% | 1.421% | 4% | 2.369% | 0.24% | 0.1421%
| 0.65 | 852.8814 | 15.8702 | 2.941% | 2.315% | 1.911% | 1.504% | 4% | 2.315% | 0.26% | 0.1504%
| 0.7 | 918.4877 | 18.4057 | 2.812% | 2.271% | 1.968% | 1.59% | 4% | 2.271% | 0.28% | 0.159%
| 0.75 | 984.094 | 21.129 | 2.707% | 2.236% | 2.031% | 1.677% | 4% | 2.236% | 0.3% | 0.1677%
| 0.8 | 1049.7002 | 24.0401 | 2.622% | 2.208% | 2.098% | 1.766% | 4% | 2.208% | 0.32% | 0.1766%
| 0.85 | 1115.3065 | 27.139 | 2.551% | 2.184% | 2.169% | 1.856% | 4% | 2.184% | 0.34% | 0.1856%
| 0.9 | 1180.9128 | 30.4258 | 2.492% | 2.164% | 2.243% | 1.948% | 4% | 2.164% | 0.36% | 0.1948%
| 0.95 | 1246.519 | 33.9003 | 2.441% | 2.147% | 2.319% | 2.04% | 4% | 2.147% | 0.38% | 0.204%
| 1 | 1312.1253 | 37.5627 | 2.399% | 2.133% | 2.399% | 2.133% | 4% | 2.133% | 0.4% | 0.2133%
|  | dp bei 100% | 37.5627
Berechnung Fluid Velocity |  | Durchfluss | 1034.32 | kg/h
|  | Dichte | 1.1889 | kg/m³
|  | Volumenstrom | 0.24165 | m³/s
|  | ID | 107.3 | mm
|  | A | 0.0090425 | m²
|  | Velocity | 26.7241 | m/s
Venturi typen, Messbereiche aus Versuchen mit Luft
Sonden als Polynom aus Versuchen, X-Achse ID^2, yAchse Messbereichsendwert, Da ein Polynom vor allem im unteren ID Berech nicht gut passt, gibt's 3 Polynome für kleine (<50mm), mittlere (<200mm) und große Leitingen (>200mm)
Die Polynome wurden mit Cuveexpert ermittelt. Wenn neue Versuche vorliegen,m müssen die Polynome neu ermittelt werden
Umrechnungsfaktor Durchfluss von Standardeinheit (kg/h] auf Usereinheit
Berechnung von Druckverlusten und Expansionszahlen bei 400mbar Werte bei Sonden sind aus DF12 deltacalc bis 100mm übernommen,
bei größeren Durchmessern wurde der Druckverlust Reduziert wegen v
erkürzter Einbaulänge. Die Expansionszahl wurde ebenfalls von DF12 Übernommen. Die Polynome wurde mit CurveExpert ermittelt
Für Venturis wurde beides zunächst aus ISO-Berechnung übernommen.
Die Drucklverluste und Expansionszahlen beziehen sich auf Normluft (0°C, 101,325 kPa)